## Asset_Cal_Info
Sensor Code | startDateTime | stopDateTime | Sensor.uid | Sensor Serial Number | Calibration Cofficient Name | Calibration Cofficient Value | Notes | Source | Undefined Sensors
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_dark_counts_volume_scatter | 51 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From CP01CNSM-00003
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_scale_factor_volume_scatter | 1.814e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_dark_counts_chlorophyll_a | 52 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_scale_factor_chlorophyll_a | 0.0121 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_dark_counts_cdom | 49 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_scale_factor_cdom | 0.0907 | Multiplier [ppb counts^-1]
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_scattering_angle | 124 | Optical backscatter scattering angle [degrees].
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_measurement_wavelength | 700 | Optical backscatter measurement wavelength [nm].
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTD | 2013-06-01 |  | A00085 | 993 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_dark_counts_volume_scatter | 50 |  | From CP01CNSM-00001
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_scale_factor_volume_scatter | 1.907e-06
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_dark_counts_chlorophyll_a | 55
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_dark_counts_cdom | 49
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_scale_factor_cdom | 0.0906
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_scattering_angle | 124
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_measurement_wavelength | 700
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_angular_resolution | 1.076
FLORTD | 2013-06-01 | 2015-01-01 | A00084 | 994 | CC_depolarization_ratio | 0.039
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_dark_counts_volume_scatter | 46 |  | From CP04OSSM-00002
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_scale_factor_volume_scatter | 1.948e-06
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_scale_factor_chlorophyll_a | 0.012
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_dark_counts_cdom | 45
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_scale_factor_cdom | 0.0898
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-01-01 |  | A00084 | 994 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_dark_counts_volume_scatter | 43 |  | From GI01SUMO-00003
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_scale_factor_volume_scatter | 1.872e-06
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_dark_counts_chlorophyll_a | 47
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_scale_factor_chlorophyll_a | 0.0116
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_dark_counts_cdom | 41
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_scale_factor_cdom | 0.0734
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_scattering_angle | 124 | Constant
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2016-01-01 |  | A00084 | 994 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_dark_counts_volume_scatter | 54 |  | From CE01ISSM-00001
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_scale_factor_volume_scatter | 1.874e-06
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_dark_counts_cdom | 49
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_scale_factor_cdom | 0.0907
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_scattering_angle | 124 | Constant; updated value
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_angular_resolution | 1.076 | CHI FACTOR. Constant; updated value
FLORTD | 2012-11-07 |  | A00218 | 995 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_dark_counts_volume_scatter | 50 |  | From CE01ISSM-00003
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_scale_factor_volume_scatter | 1.794e-06
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_dark_counts_chlorophyll_a | 48
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_scale_factor_chlorophyll_a | 0.0117
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_dark_counts_cdom | 44
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_scale_factor_cdom | 0.0857
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_scattering_angle | 124
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_measurement_wavelength | 700
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_angular_resolution | 1.076
FLORTD | 2015-01-01 |  | A00218 | 995 | CC_depolarization_ratio | 0.039
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_dark_counts_volume_scatter | 48 |  | From CE01ISSM-00005
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_scale_factor_volume_scatter | 1.786e-06
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_dark_counts_chlorophyll_a | 45
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_scale_factor_chlorophyll_a | 0.0116
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_dark_counts_cdom | 45
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_scale_factor_cdom | 0.0747
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_scattering_angle | 124 | Constant; updated value
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_angular_resolution | 1.076 | Constant; updated value
FLORTD | 2016-01-01 |  | A00218 | 995 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_dark_counts_volume_scatter | 54 |  | From CE06ISSM-00001
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_scale_factor_volume_scatter | 1.878e-06
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_dark_counts_chlorophyll_a | 60
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_scale_factor_chlorophyll_a | 0.0122
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_dark_counts_cdom | 49
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_scale_factor_cdom | 0.0906
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_scattering_angle | 124 | Constant
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2012-11-07 |  | A00219 | 996 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_scale_factor_volume_scatter | 2.118e-06 |  | From GP03FLMA-00001
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_dark_counts_volume_scatter | 54
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_dark_counts_chlorophyll_a | 53
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_scale_factor_cdom | 0.0911
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_dark_counts_cdom | 50
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_scattering_angle | 124 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_measurement_wavelength | 700 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_angular_resolution | 1.076 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00155 | 1003 | CC_depolarization_ratio | 0.039 | Default value per <flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)>
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_scale_factor_volume_scatter | 2.118e-06 |  | From GP03FLMA-00001
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_dark_counts_volume_scatter | 54
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_dark_counts_chlorophyll_a | 53
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_scale_factor_cdom | 0.0911
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_dark_counts_cdom | 50
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_scattering_angle | 124 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_measurement_wavelength | 700 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_angular_resolution | 1.076 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00153 | 1004 | CC_depolarization_ratio | 0.039 | Default value per <flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)>
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_scale_factor_volume_scatter | 2.116e-06 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From GP03FLMB-00001
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_dark_counts_volume_scatter | 65 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_scale_factor_chlorophyll_a | 0.0121 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_dark_counts_chlorophyll_a | 56 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_scale_factor_cdom | 0.0905 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_dark_counts_cdom | 48 | Multiplier [ppb counts^-1]
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_scattering_angle | 124 | Default values defined in DPA: def flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_measurement_wavelength | 700 | Default values defined in DPA: def flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTD | 2013-01-01 | 2014-06-01 | A00153 | 1004 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_scale_factor_volume_scatter | 1.675e-06 |  | From GP03FLMB-00003
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_dark_counts_volume_scatter | 51
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_dark_counts_chlorophyll_a | 55
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_scale_factor_cdom | 0.0904
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_dark_counts_cdom | 49
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-01-01 |  | A00153 | 1004 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_scale_factor_volume_scatter | 2.118e-06 |  | From GP03FLMA-00001
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_dark_counts_volume_scatter | 54
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_dark_counts_chlorophyll_a | 53
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_scale_factor_cdom | 0.0911
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_dark_counts_cdom | 50
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_scattering_angle | 124 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_measurement_wavelength | 700 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_angular_resolution | 1.076 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2012-11-07 |  | A00154 | 1005 | CC_depolarization_ratio | 0.039 | Default value per <flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)>
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_dark_counts_volume_scatter | 51 | Measured signal output of scatterometer in clean water with black tape over the detector [counts] | From CE06ISSP-00004
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_scale_factor_volume_scatter | 3.396e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_dark_counts_chlorophyll_a | 43 | Measured signal output of fluorometer in clean water with black tape over the detector [counts]
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_scale_factor_chlorophyll_a | 0.0112 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_dark_counts_cdom | 44 | Measured signal output of fluorometer in clean water with black tape over the detector [counts]
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_scale_factor_cdom | 0.0864 | Multiplier [ppb counts^-1]
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_scattering_angle | 124 | Optical backscatter scattering angle [degrees].
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_measurement_wavelength | 700 | Optical backscatter measurement wavelength [nm].
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_angular_resolution | 1.076 | X (Chi) factor (this is not really "angular resolution")
FLORTD | 2016-01-01 |  | N00593 | 1084 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_dark_counts_volume_scatter | 52 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From GI01SUMO-00001
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_scale_factor_volume_scatter | 1.82e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_dark_counts_chlorophyll_a | 53 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_scale_factor_chlorophyll_a | 0.0123 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_dark_counts_cdom | 49 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_scale_factor_cdom | 0.0905 | Multiplier [ppb counts^-1]
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_scattering_angle | 124 | Default values defined in DPA: def flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_measurement_wavelength | 700 | Default values defined in DPA: def flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTD | 2013-09-09 |  | A00409 | 1102 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_dark_counts_volume_scatter | 47 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From CP01ISSM-00004
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_scale_factor_volume_scatter | 1.803e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_dark_counts_chlorophyll_a | 53 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_scale_factor_chlorophyll_a | 0.0117 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_dark_counts_cdom | 50 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_scale_factor_cdom | 0.0844 | Multiplier [ppb counts^-1]
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_scattering_angle | 124 | Default values defined in DPA: def flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_measurement_wavelength | 700 | Default values defined in DPA: def flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTD | 2016-01-01 |  | A00409 | 1102 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_dark_counts_volume_scatter | 51 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From GI01SUMO-00001
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_scale_factor_volume_scatter | 1.675e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_dark_counts_chlorophyll_a | 55 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_scale_factor_chlorophyll_a | 0.0121 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_dark_counts_cdom | 49 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_scale_factor_cdom | 0.0904 | Multiplier [ppb counts^-1]
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_scattering_angle | 124 | Default values defined in DPA: def flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_measurement_wavelength | 700 | Default values defined in DPA: def flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTD | 2014-06-01 | 2015-01-01 | A00419 | 1104 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_dark_counts_volume_scatter | 59 |  | From GA01SUMO-00001
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_scale_factor_volume_scatter | 1.684e-06
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_dark_counts_chlorophyll_a | 58
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_scale_factor_chlorophyll_a | 0.0122
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_dark_counts_cdom | 49
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_scale_factor_cdom | 0.0901
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_scattering_angle | 124 | Constant
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2013-09-17 |  | A00420 | 1105 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_dark_counts_volume_scatter | 37 |  | From GI01SUMO-00003
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_scale_factor_volume_scatter | 1.745e-06
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_scale_factor_chlorophyll_a | 0.0116
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_dark_counts_cdom | 42
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_scale_factor_cdom | 0.0739
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_scattering_angle | 124
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_measurement_wavelength | 700
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_angular_resolution | 1.076
FLORTD | 2013-09-17 |  | A00424 | 1106 | CC_depolarization_ratio | 0.039
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_scale_factor_volume_scatter | 1.695e-06 |  | From GA01SUMO-00001
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_dark_counts_volume_scatter | 52
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_scale_factor_chlorophyll_a | 0.0122
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_dark_counts_chlorophyll_a | 59
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_scale_factor_cdom | 0.0903
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_dark_counts_cdom | 47
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_scattering_angle | 124 | Constant
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2013-09-17 |  | A00421 | 1107 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_scale_factor_volume_scatter | 1.654e-06 |  | From GP03FLMB-00004
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_dark_counts_volume_scatter | 48
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_scale_factor_chlorophyll_a | 0.0113
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_dark_counts_chlorophyll_a | 52
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_scale_factor_cdom | 0.0775
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_dark_counts_cdom | 45
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_scattering_angle | 124 | Constant
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2016-01-01 |  | A00421 | 1107 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_scale_factor_volume_scatter | 1.777e-06 |  | From GP03FLMA-00002
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_dark_counts_volume_scatter | 52
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_dark_counts_chlorophyll_a | 54
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_scale_factor_cdom | 0.0903
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_dark_counts_cdom | 49
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_scattering_angle | 124 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_measurement_wavelength | 700 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_angular_resolution | 1.076 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2013-10-09 |  | A00470 | 1116 | CC_depolarization_ratio | 0.039 | Default value per <flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)>
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_scale_factor_volume_scatter | 1.816e-06 |  | From GP03FLMA-00004
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_dark_counts_volume_scatter | 49
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_scale_factor_chlorophyll_a | 0.0118
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_dark_counts_chlorophyll_a | 49
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_scale_factor_cdom | 0.0902
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_dark_counts_cdom | 46
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_scattering_angle | 124 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_measurement_wavelength | 700 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_angular_resolution | 1.076 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2016-01-01 |  | A00470 | 1116 | CC_depolarization_ratio | 0.039 | Default value per <flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)>
2016-01-01
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_dark_counts_volume_scatter | 52 |  | From CE06ISSM-00001
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scale_factor_volume_scatter | 1.74e-06
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scale_factor_chlorophyll_a | 0.0117
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_dark_counts_cdom | 45
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scale_factor_cdom | 0.085
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scattering_angle | 124 | Constant; updated value
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_angular_resolution | 1.076 | Constant; updated value
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_dark_counts_volume_scatter | 52 |  | From CE06ISSM_00003
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scale_factor_volume_scatter | 1.74e-06
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scale_factor_chlorophyll_a | 0.0117
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_dark_counts_cdom | 45
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scale_factor_cdom | 0.085
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_scattering_angle | 124 | Constant; updated value
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_angular_resolution | 1.076 | Constant; updated value
FLORTD | 2013-10-11 |  | A00474 | 1121 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_scale_factor_volume_scatter | 1.819e-06 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From GI03FLMA-00001
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_dark_counts_volume_scatter | 52 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_scale_factor_chlorophyll_a | 0.0117 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_dark_counts_chlorophyll_a | 52 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_scale_factor_cdom | 0.0903 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_dark_counts_cdom | 45 | Multiplier [ppb counts^-1]
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_scattering_angle | 124 | Default values defined in DPA: def flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_measurement_wavelength | 700 | Default values defined in DPA: def flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTD | 2013-10-09 |  | A00466 | 1122 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_scale_factor_volume_scatter | 1.623e-06 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From GI03FLMA-00003
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_dark_counts_volume_scatter | 50 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_scale_factor_chlorophyll_a | 0.0107 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_dark_counts_chlorophyll_a | 50 | Multiplier [ug L^-1 counts^-1]
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_scale_factor_cdom | 0.0733 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_dark_counts_cdom | 46 | Multiplier [ppb counts^-1]
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_scattering_angle | 124 | Default values defined in DPA: def flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_measurement_wavelength | 700 | Default values defined in DPA: def flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTD | 2016-01-01 |  | A00466 | 1122 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_dark_counts_volume_scatter | 51 |  | From CE02SHSM-00001
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scale_factor_volume_scatter | 1.727e-06
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_dark_counts_chlorophyll_a | 51
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scale_factor_chlorophyll_a | 0.0119
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_dark_counts_cdom | 46
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scale_factor_cdom | 0.098
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scattering_angle | 124 | Constant
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2016-01-01 |  | A00475 | 1123 | CC_dark_counts_volume_scatter | 50 |  | From CE02SHSM-00003
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scale_factor_volume_scatter | 1.858e-06
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_dark_counts_chlorophyll_a | 48
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scale_factor_chlorophyll_a | 0.0115
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_dark_counts_cdom | 46
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scale_factor_cdom | 0.0734
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_scattering_angle | 124 | Constant
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2013-10-11 |  | A00475 | 1123 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_scale_factor_volume_scatter | 1.71e-06 |  | From GI03FLMB-00002
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_dark_counts_volume_scatter | 50
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_scale_factor_chlorophyll_a | 0.012
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_dark_counts_chlorophyll_a | 56
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_scale_factor_cdom | 0.0902
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_dark_counts_cdom | 44
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_scattering_angle | 124 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_measurement_wavelength | 700 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_angular_resolution | 1.076 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
FLORTD | 2013-10-09 |  | A00467 | 1124 | CC_depolarization_ratio | 0.039 | Default value per <flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)>
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_scattering_angle | 124 |  | From RS03AXPS
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_depolarization_ratio | 0.039
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_measurement_wavelength | 700
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_angular_resolution | 1.076
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_dark_counts_volume_scatter | 56 | counts
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_scale_factor_volume_scatter | 1.797e-06 | (m^-1·sr^-1)/counts
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_dark_counts_chlorophyll_a | 56 | counts
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_scale_factor_chlorophyll_a | 0.0122 | µg/l/count
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_dark_counts_cdom | 48 | counts
FLORTD | 2014-06-01 |  | ATAPL-58322-00002 | 1129 | CC_scale_factor_cdom | 0.0904 | ppb/count
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_depolarization_ratio | 0.039 |  | From CE04OSPS-00001
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_scattering_angle | 124
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_measurement_wavelength | 700
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_angular_resolution | 1.076
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_dark_counts_volume_scatter | 52
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_scale_factor_volume_scatter | 1.896e-06
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_dark_counts_chlorophyll_a | 53
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_scale_factor_chlorophyll_a | 0.0073
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_dark_counts_cdom | 43
FLORTD | 2014-02-18 |  | A00695 | 1150 | CC_scale_factor_cdom | 0.0902
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_dark_counts_volume_scatter | 53 |  | From CE04OSSM-00001
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_scale_factor_volume_scatter | 1.877e-06
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_dark_counts_chlorophyll_a | 48
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_scale_factor_chlorophyll_a | 0.0073
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_dark_counts_cdom | 41
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_scale_factor_cdom | 0.0902
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-02-18 |  | A00696 | 1151 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_dark_counts_volume_scatter | 50 |  | From CE07SHSM-00001
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_scale_factor_volume_scatter | 1.833e-06
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_dark_counts_chlorophyll_a | 46
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_scale_factor_chlorophyll_a | 0.0073
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_dark_counts_cdom | 43
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_scale_factor_cdom | 0.0905
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-01-03 |  | A00614 | 1152 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_dark_counts_volume_scatter | 50 |  | From CE02SHSM-00003
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_scale_factor_volume_scatter | 1.935e-06
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_dark_counts_chlorophyll_a | 45
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_scale_factor_chlorophyll_a | 0.007
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_dark_counts_cdom | 45
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_scale_factor_cdom | 0.0811
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_scattering_angle | 124 | Constant
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2016-01-01 |  | A00614 | 1152 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_dark_counts_volume_scatter | 49 |  | From CE01ISSM-00002
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_scale_factor_volume_scatter | 1.837e-06
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_scale_factor_chlorophyll_a | 0.0073
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_dark_counts_cdom | 43
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_scale_factor_cdom | 0.0896
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-01-03 |  | A00615 | 1153 | CC_depolarization_ratio | 0.039 | Constant
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_scattering_angle
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_measurement_wavelength
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_angular_resolution
[9 empty] | FLORTD | 2014-01-03 | A00616 | 1154 | CC_depolarization_ratio
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_dark_counts_volume_scatter | 50 |  | From CE09OSSM-00001
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_scale_factor_volume_scatter | 1.84e-06
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_scale_factor_chlorophyll_a | 0.0073
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_dark_counts_cdom | 46
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_scale_factor_cdom | 0.0903
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-01-03 |  | A00617 | 1155 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_dark_counts_volume_scatter | 50 |  | From CE09OSSM-00003
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_scale_factor_volume_scatter | 1.913e-06
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_dark_counts_chlorophyll_a | 52
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_scale_factor_chlorophyll_a | 0.0071
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_dark_counts_cdom | 50
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_scale_factor_cdom | 0.0877
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_scattering_angle | 124 | Constant; not 117
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_angular_resolution | 1.076 | Constant; not 1.08
FLORTD | 2016-01-01 |  | A00617 | 1155 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_dark_counts_volume_scatter | 49 |  | From CE06ISSM-00002
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_scale_factor_volume_scatter | 1.915e-06
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_dark_counts_chlorophyll_a | 56
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_dark_counts_cdom | 50
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_scale_factor_cdom | 0.0823
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_scattering_angle | 124 | Constant; updated value
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_angular_resolution | 1.076 | Constant; updated value
FLORTD | 2014-12-13 |  | A01229 | 1197 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_dark_counts_volume_scatter | 50 |  | From CE06ISSM-00004
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_scale_factor_volume_scatter | 1.835e-06
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_dark_counts_chlorophyll_a | 48
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_dark_counts_cdom | 45
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_scale_factor_cdom | 0.0904
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_scattering_angle | 124 | Constant
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2016-01-01 |  | A01229 | 1197 | CC_depolarization_ratio | 0.039 | Constant
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_dark_counts_volume_scatter | 48 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From CE01ISSP-00005
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_scale_factor_volume_scatter | 2.917e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_dark_counts_chlorophyll_a | 48 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_scale_factor_chlorophyll_a | 0.0118 | Multiplier [ug L^-1 counts^-1]
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_dark_counts_cdom | 45 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_scale_factor_cdom | 0.0832 | Multiplier [ppb counts^-1]
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_scattering_angle | 124 | Optical backscatter scattering angle [degrees].
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_measurement_wavelength | 700 | Optical backscatter measurement wavelength [nm].
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTJ | 2015-01-01 |  | N00585 | 1206 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_dark_counts_volume_scatter | 46 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From CE01ISSP-00004
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_scale_factor_volume_scatter | 3.016e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_dark_counts_chlorophyll_a | 50 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_scale_factor_chlorophyll_a | 0.0121 | Multiplier [ug L^-1 counts^-1]
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_dark_counts_cdom | 41 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_scale_factor_cdom | 0.0904 | Multiplier [ppb counts^-1]
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_scattering_angle | 124 | Optical backscatter scattering angle [degrees].
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_measurement_wavelength | 700 | Optical backscatter measurement wavelength [nm].
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTJ | 2015-01-01 |  | N00577 | 1207 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_dark_counts_volume_scatter | 50 | Measured signal output of fluormeter in clean water with black tape over the detector [counts] | From CP03ISSM-00003
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_scale_factor_volume_scatter | 3.092e-06 | Multiplier [m^-1 sr^-1 counts^-1]
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_dark_counts_chlorophyll_a | 47 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_scale_factor_chlorophyll_a | 0.0121 | Multiplier [ug L^-1 counts^-1]
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_dark_counts_cdom | 40 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_scale_factor_cdom | 0.0903 | Multiplier [ppb counts^-1]
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_scattering_angle | 124 | Optical backscatter scattering angle [degrees].
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_measurement_wavelength | 700 | Optical backscatter measurement wavelength [nm].
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
FLORTJ | 2015-01-01 |  | N00569 | 1208 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_dark_counts_volume_scatter | 43 |  | From GS01SUMO-00001
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_scale_factor_volume_scatter | 1.883e-06
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_dark_counts_chlorophyll_a | 51
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_scale_factor_chlorophyll_a | 0.0122
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_dark_counts_cdom | 42
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_scale_factor_cdom | 0.0903
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-03-14 |  | A00992 | 1212 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_dark_counts_volume_scatter | 50 |  | From CP01CNSM-00005
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_scale_factor_volume_scatter | 1.83e-06
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_dark_counts_chlorophyll_a | 49
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_dark_counts_cdom | 45
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_scale_factor_cdom | 0.0908
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-06-27 |  | A00888 | 1213 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_dark_counts_volume_scatter | 51 |  | From GA01SUMO-00001
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_scale_factor_volume_scatter | 1.851e-06
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_dark_counts_chlorophyll_a | 52
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_dark_counts_cdom | 43
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_scale_factor_cdom | 0.0908
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-06-27 |  | A00889 | 1214 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_scale_factor_volume_scatter | 1.87e-06 |  | From GA03FLMA-00001
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_dark_counts_volume_scatter | 51
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_dark_counts_chlorophyll_a | 52
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_scale_factor_cdom | 0.0906
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_dark_counts_cdom | 45
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-01-01 |  | A01496 | 1215 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_scale_factor_volume_scatter | 1.929e-06 |  | From GS03FLMA-00001
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_dark_counts_volume_scatter | 48
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_dark_counts_chlorophyll_a | 48
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_scale_factor_cdom | 0.0905
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_dark_counts_cdom | 42
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-09-01 |  | A01497 | 1216 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_scale_factor_volume_scatter | 1.929e-06 |  | From GP03FLMA-00003
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_dark_counts_volume_scatter | 48
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_dark_counts_chlorophyll_a | 48
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_scale_factor_cdom | 0.0905
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_dark_counts_cdom | 42
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-01-01 |  | A01497 | 1216 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_scale_factor_volume_scatter | 1.872e-06 |  | From GS03FLMB-00001
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_dark_counts_volume_scatter | 45
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_dark_counts_chlorophyll_a | 51
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_scale_factor_cdom | 0.0909
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_dark_counts_cdom | 43
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-09-01 |  | A01498 | 1217 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_dark_counts_volume_scatter | 50 |  | From GS01SUMO-00001
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_scale_factor_volume_scatter | 1.892e-06
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_dark_counts_cdom | 46
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_scale_factor_cdom | 0.0908
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-03-14 |  | A00993 | 1218 | CC_depolarization_ratio | 0.039 | Constant
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_scattering_angle
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_angular_resolution
[9 empty] | FLORTD | 2015-09-04 | A01499 | 1219 | CC_depolarization_ratio
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_dark_counts_volume_scatter | 51 |  | From CP04OSSM-00001
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_scale_factor_volume_scatter | 1.885e-06
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_dark_counts_chlorophyll_a | 53
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_scale_factor_chlorophyll_a | 0.0122
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_dark_counts_cdom | 43
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_scale_factor_cdom | 0.0906
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-03-14 | 2015-05-31 13:00 | A00994 | 1221 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_dark_counts_volume_scatter | 51 |  | From CP01CNSM-00004
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_scale_factor_volume_scatter | 1.885e-06
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_dark_counts_chlorophyll_a | 53
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_scale_factor_chlorophyll_a | 0.0122
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_dark_counts_cdom | 43
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_scale_factor_cdom | 0.0906
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-06-01 |  | A00994 | 1221 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_scale_factor_volume_scatter | 2.04e-06 |  | From GI03FLMA-00003
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_dark_counts_volume_scatter | 51
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_scale_factor_chlorophyll_a | 0.011
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_dark_counts_chlorophyll_a | 55
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_scale_factor_cdom | 0.0797
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_dark_counts_cdom | 50
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_scattering_angle | 124 | Constant
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2016-01-01 |  | A00994 | 1221 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_dark_counts_volume_scatter | 64 |  | From CP03ISSM-00002, CP04OSSM-00004
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_scale_factor_volume_scatter | 1.953e-06
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_dark_counts_chlorophyll_a | 47
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_dark_counts_cdom | 44
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_scale_factor_cdom | 0.0904
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_scattering_angle | 124 | Constant
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2014-08-21 |  | A00995 | 1222 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_depolarization_ratio | 0.039 |  | From CE04OSPS-00002
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_scattering_angle | 124
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_measurement_wavelength | 700
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_angular_resolution | 1.076
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_dark_counts_volume_scatter | 48
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_scale_factor_volume_scatter | 1.842e-06
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_dark_counts_chlorophyll_a | 49
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_dark_counts_cdom | 51
FLORTD | 2015-06-03 |  | A01696 | 1289 | CC_scale_factor_cdom | 0.0905
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_scattering_angle
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_angular_resolution
[9 empty] | FLORTD | 2015-07-16 | A01684 | 1290 | CC_depolarization_ratio
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_dark_counts_volume_scatter | 48 |  | From CE09OSSM-00002
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_scale_factor_volume_scatter | 1.839e-06
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_dark_counts_chlorophyll_a | 48
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_dark_counts_cdom | 45
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_scale_factor_cdom | 0.0907
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_scattering_angle | 124 | Constant; not 117
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_angular_resolution | 1.076 | Constant; not 1.08
FLORTD | 2015-07-16 |  | A01685 | 1291 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_depolarization_ratio | 0.039 |  | From RS01SBPS
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_scattering_angle | 124
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_measurement_wavelength | 700
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_angular_resolution | 1.076
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_dark_counts_volume_scatter | 51
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_scale_factor_volume_scatter | 1.842e-06
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_dark_counts_chlorophyll_a | 52
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_dark_counts_cdom | 45
FLORTD | 2013-02-13 |  | ATAPL-58322-00009 | 1292 | CC_scale_factor_cdom | 0.0906
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_scattering_angle | 124 |  | From RS03AXPS
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_depolarization_ratio | 0.039
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_measurement_wavelength | 700
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_angular_resolution | 1.076
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_dark_counts_volume_scatter | 52 | counts
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_scale_factor_volume_scatter | 1.814e-06 | (m^-1·sr^-1)/counts
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_dark_counts_chlorophyll_a | 51 | counts
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_scale_factor_chlorophyll_a | 0.0121 | µg/l/count
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_dark_counts_cdom | 43 | counts
FLORTD | 2015-01-01 |  | ATAPL-58322-00010 | 1293 | CC_scale_factor_cdom | 0.0906 | ppb/count
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_scattering_angle
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_angular_resolution
[9 empty] | FLORTD | 2015-06-01 | N00724 | 1294 | CC_depolarization_ratio
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_dark_counts_volume_scatter | 41 |  | From GI01SUMO-00002
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_scale_factor_volume_scatter | 1.757e-06
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_dark_counts_chlorophyll_a | 56
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_scale_factor_chlorophyll_a | 0.0122
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_dark_counts_cdom | 46
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_scale_factor_cdom | 0.0905
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-04-14 |  | A01340 | 1295 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_dark_counts_volume_scatter | 57 |  | From GI01SUMO-00002
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_scale_factor_volume_scatter | 1.819e-06
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_dark_counts_chlorophyll_a | 51
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_dark_counts_cdom | 46
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_scale_factor_cdom | 0.0906
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-04-14 |  | A01341 | 1296 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_scale_factor_volume_scatter | 1.77e-06 |  | From GI03FLMA-00002
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_dark_counts_volume_scatter | 46
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_dark_counts_chlorophyll_a | 46
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_scale_factor_cdom | 0.0908
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_dark_counts_cdom | 43
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-05-14 |  | A01424 | 1298 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_scale_factor_volume_scatter | 1.785e-06 |  | From GI03FLMB-00002
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_dark_counts_volume_scatter | 50
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_scale_factor_cdom | 0.0906
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_dark_counts_cdom | 44
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-05-14 |  | A01425 | 1299 | CC_depolarization_ratio | 0.039 | Constant
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_scattering_angle
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_angular_resolution
[9 empty] | FLORTD | 2015-05-18 | A01429 | 1300 | CC_depolarization_ratio
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_scattering_angle
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_angular_resolution
[9 empty] | FLORTD | 2015-05-18 | A01430 | 1301 | CC_depolarization_ratio
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_dark_counts_volume_scatter | 54 |  | From CE06SHSM-00002
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_scale_factor_volume_scatter | 1.819e-06
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_dark_counts_chlorophyll_a | 52
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_dark_counts_cdom | 44
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_scale_factor_cdom | 0.0907
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_scattering_angle | 124 | Constant; not 117
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_angular_resolution | 1.076 | Constant; not 1.08
FLORTD | 2015-05-18 |  | A01431 | 1302 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_dark_counts_volume_scatter | 50 |  | From CE02SHSM-00002
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_scale_factor_volume_scatter | 1.85e-06
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_dark_counts_chlorophyll_a | 51
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_dark_counts_cdom | 46
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_scale_factor_cdom | 0.0906
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-07-16 |  | A01686 | 1303 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_dark_counts_volume_scatter | 50 |  | From GA01SUMO-00002
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_scale_factor_volume_scatter | 1.836e-06
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_dark_counts_chlorophyll_a | 52
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_dark_counts_cdom | 43
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_scale_factor_cdom | 0.0921
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-07-15 |  | A01630 | 1317 | CC_depolarization_ratio | 0.039 | Constant
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_dark_counts_volume_scatter | 51 |  | From GA01SUMO-00002
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_scale_factor_volume_scatter | 1.861e-06
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_dark_counts_chlorophyll_a | 51
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_dark_counts_cdom | 45
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_scale_factor_cdom | 0.0901
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_scattering_angle | 124 | Constant
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_measurement_wavelength | 700 | Constant
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_angular_resolution | 1.076 | Constant
FLORTD | 2015-07-15 |  | A01631 | 1318 | CC_depolarization_ratio | 0.039 | Constant
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_scattering_angle
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_angular_resolution
[9 empty] | FLORTD | 2015-07-15 | A01632 | 1319 | CC_depolarization_ratio
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_scattering_angle
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_angular_resolution
[9 empty] | FLORTD | 2015-07-15 | A01633 | 1320 | CC_depolarization_ratio
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_scattering_angle
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_angular_resolution
[9 empty] | FLORTD | 2015-04-06 | A01651 | 1321 | CC_depolarization_ratio
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_scattering_angle
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_angular_resolution
[9 empty] | FLORTD | 2015-04-06 | A01652 | 1322 | CC_depolarization_ratio
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_scattering_angle
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_measurement_wavelength
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_angular_resolution
[9 empty] | FLORTD | 2013-01-25 | A10653 | 1323 | CC_depolarization_ratio
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_dark_counts_volume_scatter | 49 |  | From CP03ISSM-00003
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_scale_factor_volume_scatter | 1.931e-06
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_dark_counts_chlorophyll_a | 50
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_scale_factor_chlorophyll_a | 0.0121
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_dark_counts_cdom | 41
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_scale_factor_cdom | 0.0906
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_scattering_angle | 124
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_measurement_wavelength | 700
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_angular_resolution | 1.076
FLORTD | 2015-04-06 |  | A01654 | 1324 | CC_depolarization_ratio | 0.039
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_dark_counts_volume_scatter
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_scale_factor_volume_scatter
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_dark_counts_chlorophyll_a
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_scale_factor_chlorophyll_a
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_dark_counts_cdom
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_scale_factor_cdom
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_scattering_angle
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_measurement_wavelength
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_angular_resolution
[9 empty] | FLORTD | 2015-06-01 | OL000031 | 1326 | CC_depolarization_ratio